## transmission&scatter
date: | 2014-07-08 |  |  |  | Tref [°C] | 20
linac: | U7 |  |  |  | T0 [°C] | -273.15
field size: | 25x25 |  |  |  | DOSE 1cal | 59.16 | [mC/Gy]
chamber: | NE2571, sn 3593 |  |  |  | pref [hPa] | 1013.25
electrometer: | PTW Webline, SN176, -250 V
phantoms: | RW3-stack
PMMA-cylinder (torentjesfantoom, 3 cm @ 10 cm)
depth: | 10 cm
ssd: | 100 cm
Stealth-device: | A
measurement is average over 10 samples
dose: | 100 ME
puls rate: | 400 | s-1
time: | 11.25 | s
pulses: | 4500
phantom | MV | Stealth | m [nC] | Tphantom [°C] | m22 [nC] | Sc/Scp | scatter contribution
PMMA | 6 | TRUE | 11.11 | 22.9 | 11.2199 | 0.69885 | -0.2462%
PMMA | 10 | TRUE | 12.91 | 22.9 | 13.0377 | 0.75807 | 0.15%
PMMA | 10 | FALSE | 13.05 | 22.9 | 13.1791 | 0.75693
PMMA | 6 | FALSE | 11.32 | 22.9 | 11.432 | 0.70058
RW3 | 6 | FALSE | 16.12 | 23.6 | 16.318
RW3 | 10 | FALSE | 17.2 | 23.6 | 17.4112
RW3 | 10 | TRUE | 16.99 | 23.6 | 17.1986
RW3 | 6 | TRUE | 15.86 | 23.6 | 16.0548
| transmission |  |  |  |  | scatter
| MV | PMMA | RW3 |  |  | MV | open | Stealth | difference
| 6 | 98.14% | 98.39% |  |  | 6 | 0.70058 | 0.69885 | -0.2462%
| 10 | 98.93% | 98.78% |  |  | 10 | 0.75693 | 0.75807 | 0.15%
conclusions
transmission approximately 98.4% for 6 MV and 98.8% for 10 MV
scatter contribution below detection limit

## linearity
date: | 2014-07-08 |  | VDOSE 1 | V-1 | mon [nC] | Stealth [μGy] | Stealth [nC] | Qrel-1 [*10-5] | Q-1 [*10-5] | avg Q-1
linac: | U7 |  | 24 | 0.041667 | 34.5 | 163.3 | 9660.828 | 10.3391 | 10.3511 | 10.3676 | 1.9322
field size: | 25x25 |  | 0.09818% |  | 34.55 | 163.1 | 9648.996 | 10.3668 | 10.3638
chamber: | NE2571, sn 3593 |  | 0.1412% |  | 34.56 | 162.8 | 9631.248 | 10.3889 | 10.3829
electrometer: | PTW Webline, SN176, -250 V |  | 0.1745% |  | 34.51 | 162.8 | 9631.248 | 10.3739 | 10.3829
phantoms: | - |  | 34.54 |  | 34.58 | 163.2 | 9654.912 | 10.3694 | 10.3574
depth: | - |  | 50 | 0.02 | 34.57 | 249 | 14730.84 | 6.7944 | 6.7885 | 6.7967 | 1.2667
ssd: | - |  | 0.0969% |  | 34.49 | 248.2 | 14683.512 | 6.8005 | 6.8104
Stealth-device: | A |  | 0.1393% |  | 34.53 | 248.5 | 14701.26 | 6.8002 | 6.8021
electrometer: | DOSE 1 |  | 0.0564% |  | 34.57 | 249 | 14730.84 | 6.7944 | 6.7885
temperature: | 23.6 |  | 34.538 |  | 34.53 | 248.8 | 14719.008 | 6.792 | 6.7939
energy: | 6 MV |  | 99 | 0.010101 | 34.54 | 294.8 | 17440.368 | 5.7338 | 5.7338 | 5.7381 | 1.0694
dose: | 100 ME |  | 0.05416% |  | 34.54 | 294.4 | 17416.704 | 5.7416 | 5.7416
|  |  | 0.06073% |  | 34.56 | 294.6 | 17428.536 | 5.741 | 5.7377
cpp [nC]: | 4.1346 |  | 0.0564% |  | 34.55 | 294.7 | 17434.452 | 5.7374 | 5.7358
|  |  | 34.54 |  | 34.51 | 294.4 | 17416.704 | 5.7366 | 5.7416
Two-voltage, linear |  |  | 199 | 0.0050251 | 34.54 | 308.8 | 18268.608 | 5.4739 | 5.4739 | 5.4785 | 1.021
ks @ 400 V | V2/V1 | M2/M1 | 0.01294% |  | 34.55 | 308.7 | 18262.692 | 5.4772 | 5.4756
1.0212 | 2.0101 | 1.021 | 0.07806% |  | 34.55 | 308.6 | 18256.776 | 5.479 | 5.4774
1.0234 | 4.0404 | 1.0694 | 0.0865% |  | 34.55 | 308.4 | 18244.944 | 5.4826 | 5.481
(non-lin) | 8 | 1.2667 | 34.548 |  | 34.55 | 308.2 | 18233.112 | 5.4861 | 5.4845
Two-voltage, Boag |  | u | 400 | 0.0025 | 34.48 | 314.5 | 18605.82 | 5.3653 | 5.3747 | 5.3658 | 1
1.0215 | 3.7175e-10 | 0.043379 | 0.1032% |  | 34.55 | 315.5 | 18664.98 | 5.3592 | 5.3576
1.0241 | -7.3537e-08 | 0.04854 | 0.1217% |  | 34.56 | 315.1 | 18641.316 | 5.3675 | 5.3644
1.0435 | 9.0865e-09 | 0.088268 | 0.06839% |  | 34.53 | 314.8 | 18623.568 | 5.368 | 5.3695
|  |  | 34.538 |  | 34.57 | 315.2 | 18647.232 | 5.3674 | 5.3627
|  |  | 500 | 0.002 | 34.52 | 316.6 | 18730.056 | 5.3359 | 5.339 | 5.3407 | 0.99532
|  |  | 0.04844% |  | 34.54 | 316.5 | 18724.14 | 5.3407 | 5.3407
|  |  | 0.02234% |  | 34.55 | 316.5 | 18724.14 | 5.3422 | 5.3407
|  |  | 0.06836% |  | 34.56 | 316.5 | 18724.14 | 5.3438 | 5.3407
|  |  | 34.546 |  | 34.56 | 316.4 | 18718.224 | 5.3455 | 5.3424
conclusion
Significant recombination, no gas amplification up to 500 V

## field size
date: | 2014-07-14 | VDOSE 1 |  |  | -400 V |  | VDOSE 1 |  |  | -20 V
linac: | U7 |  | charge/area | field | Stealth [μGy] | Stealth [nC] |  | charge/area | field | Stealth [μGy] | Stealth [nC]
field size: | var | 18634.2168 | 29.8147 | 25 | 314.6 | 18611.736 | 8336.8272 | 13.3389 | 25 | 140.7 | 8323.812
chamber: | - | 0.09091% |  |  | 315 | 18635.4 | 0.1365% |  |  | 140.9 | 8335.644
electrometer: | - |  |  |  | 315.2 | 18647.232 |  |  |  | 141.2 | 8353.392
phantoms: | - |  |  |  | 314.8 | 18623.568 |  |  |  | 141 | 8341.56
depth: | - |  |  |  | 315.3 | 18653.148 |  |  |  | 140.8 | 8329.728
ssd: | - | 7132.3296 | 31.6992 | 15 | 120.5 | 7128.78 | 3822.6826 | 16.9897 | 15 | 64.63 | 3823.5108
Stealth-device: | A | 0.04543% |  |  | 120.6 | 7134.696 | 0.0964% |  |  | 64.54 | 3818.1864
electrometer: | DOSE 1 |  |  |  | 120.6 | 7134.696 |  |  |  | 64.61 | 3822.3276
temperature: | 23.6 |  |  |  | 120.6 | 7134.696 |  |  |  | 64.71 | 3828.2436
energy: | 6 MV |  |  |  | 120.5 | 7128.78 |  |  |  | 64.59 | 3821.1444
dose: | 100 ME | 3166.4798 | 31.6648 | 10 | 53.46 | 3162.6936 | 1919.0321 | 19.1903 | 10 | 32.45 | 1919.742
|  | 0.111% |  |  | 53.5 | 3165.06 | 0.06683% |  |  | 32.46 | 1920.3336
[5 empty] | 53.53 | 3166.8348 |  |  |  | 32.45 | 1919.742
[5 empty] | 53.51 | 3165.6516 |  |  |  | 32.42 | 1917.9672
[5 empty] | 53.62 | 3172.1592 |  |  |  | 32.41 | 1917.3756
|  | 724.355 | 28.9742 | 5 | 12.24 | 724.1184 | 527.3996 | 21.096 | 5 | 8.929 | 528.2396
|  | 0.09312% |  |  | 12.25 | 724.71 | 0.1038% |  |  | 8.91 | 527.1156
[5 empty] | 12.24 | 724.1184 |  |  |  | 8.91 | 527.1156
[5 empty] | 12.23 | 723.5268 |  |  |  | 8.919 | 527.648
[5 empty] | 12.26 | 725.3016 |  |  |  | 8.906 | 526.879
|  | 258.3636 | 28.7071 | 3 | 4.366 | 258.2926 | 208.3734 | 23.1526 | 3 | 3.522 | 208.3615
|  | 0.0816% |  |  | 4.364 | 258.1742 | 0.05079% |  |  | 3.52 | 208.2432
[5 empty] | 4.373 | 258.7067 |  |  |  | 3.522 | 208.3615
[5 empty] | 4.365 | 258.2334 |  |  |  | 3.522 | 208.3615
[5 empty] | 4.368 | 258.4109 |  |  |  | 3.525 | 208.539
|  | 31.7796 | 31.7796 | 1 | 0.5363 | 31.7275 | 29.8474 | 29.8474 | 1 | 0.5024 | 29.722
|  | 0.2641% |  |  | 0.5365 | 31.7393 | 0.2551% |  |  | 0.5047 | 29.8581
[5 empty] | 0.5357 | 31.692 |  |  |  | 0.5058 | 29.9231
[5 empty] | 0.5387 | 31.8695 |  |  |  | 0.5052 | 29.8876
[5 empty] | 0.5387 | 31.8695 |  |  |  | 0.5045 | 29.8462
|  | 18656.6976 | 29.8507 | 25 | 315.4 | 18659.064 | 8313.1632 | 13.3011 | 25 | 140.4 | 8306.064
|  | 0.06947% |  |  | 315.4 | 18659.064 | 0.1273% |  |  | 140.4 | 8306.064
[5 empty] | 315 | 18635.4 |  |  |  | 140.4 | 8306.064
[5 empty] | 315.4 | 18659.064 |  |  |  | 140.6 | 8317.896
[5 empty] | 315.6 | 18670.896 |  |  |  | 140.8 | 8329.728
Conclusion
Charge per area roughly constant at working voltage.
At very low voltages the high mobility of the electrons compensates leads to a spectacular lower recombination for small fields.

## polarity
date: | 2014-07-08 | [8 empty] | VDOSE 1 |  | -40 |  |  |  |  | -400 |  |  |  |  | 400
linac: | U7 |  | table [cm] | RW3 [cm] | kpol- | ks | Boag-Mw/Ml | u | cpp [nC] | Qrel | mon [nC] | Stealth [μGy] | Stealth [nC] | Qrel [*10-5] | Qrel | mon [nC] | Stealth [μGy] | Stealth [nC] | Qrel [*10-5] |  | mon [nC] | Stealth [μGy] | Stealth [nC] | Qrel [*10-5]
field size: | 25x25 |  | 30.4 | 0 | 0.99977 | 1.0419 | 6.4844e-07 | 0.085029 | 2.6188 | 11769.2416 | 72.43 | 199.2 | 11784.672 | 11760.006 | 15608.3599 | 72.31 | 263.7 | 15600.492 | 15593.6745 | 15601.074 | 72.12 | 263.1 | 15564.996 | 15599.182
chamber: | NE2571, sn 3593 |  |  |  |  | (should be appr. 1.02) |  |  |  | 0.02981% | 72.29 | 198.9 | 11766.924 | 11765.0358 | 0.1033% | 72.14 | 263.5 | 15588.66 | 15618.5667 | 0.07486% | 72.19 | 263.5 | 15588.66 | 15607.749
electrometer: | PTW Webline, SN176, -250 V | [8 empty] | 0.02283% | 72.27 | 199 | 11772.84 | 11774.2084 | 0.05363% | 72.23 | 263.7 | 15600.492 | 15610.9456 | 0.08658% | 72.23 | 263.3 | 15576.828 | 15587.2657
phantoms: | - | [8 empty] | 0.01904% | 72.22 | 199 | 11772.84 | 11782.36 | 0.06402% | 72.21 | 263.7 | 15600.492 | 15615.2694 | 0.05379% | 72.15 | 263.3 | 15576.828 | 15604.5489
depth: | var | [8 empty] | 72.286 | 72.22 | 198.7 | 11755.092 | 11764.5976 | 72.242 | 72.32 | 263.9 | 15612.324 | 15603.3435 | 72.188 | 72.25 | 263.7 | 15600.492 | 15606.6242
ssd: | var |  | 10.2 | 0 | 0.99996 |  |  |  | 1.7064 | 7685.2458 | 72.22 | 129.8 | 7678.968 | 7685.1775 | 9151.0928 | 72.26 | 154.9 | 9163.884 | 9166.2175 | 9150.4127 | 72.35 | 154.6 | 9146.136 | 9137.0847
Stealth-device: | B | [8 empty] | 0.0261% | 72.08 | 129.6 | 7667.136 | 7688.2398 | 0.06606% | 72.34 | 154.7 | 9152.052 | 9144.2587 | 0.08731% | 72.24 | 154.5 | 9140.22 | 9145.0786
electrometer: | DOSE 1 | [8 empty] | 0.01723% | 72.15 | 129.7 | 7673.052 | 7686.7072 | 0.07917% | 72.29 | 154.6 | 9146.136 | 9144.6684 | 0.08433% | 72.2 | 154.5 | 9140.22 | 9150.1451
temperature: | 23.6 | °C | [7 empty] | 0.01227% | 72.2 | 129.8 | 7678.968 | 7687.3064 | 0.1009% | 72.22 | 154.6 | 9146.136 | 9153.5319 | 0.1085% | 72.28 | 154.8 | 9157.968 | 9157.7653
energy: | 6 MV | [8 empty] | 72.186 | 72.28 | 129.8 | 7678.968 | 7678.798 | 72.286 | 72.32 | 154.7 | 9152.052 | 9146.7875 | 72.254 | 72.2 | 154.7 | 9152.052 | 9161.99
dose: | 100 | ME | -30 | 0 | 0.99906 |  |  |  | 0.87507 | 3933.7567 | 72.33 | 66.562 | 3937.8079 | 3934.9987 | 4296.7063 | 72.2 | 72.91 | 4313.3556 | 4318.0394 | 4288.642 | 72.38 | 72.56 | 4292.6496 | 4286.624
gantry: | 180 | ° | [7 empty] | 0.03228% | 72.46 | 66.57 | 3938.2812 | 3928.411 | 0.07388% | 72.33 | 72.6 | 4295.016 | 4291.9519 | 0.1068% | 72.21 | 72.48 | 4287.9168 | 4291.9785
[10 empty] | 0.03068% | 72.21 | 66.38 | 3927.0408 | 3930.7606 | 0.2369% | 72.3 | 72.5 | 4289.1 | 4287.8186 | 0.0776% | 72.32 | 72.57 | 4293.2412 | 4290.7716
Vw/Vl | 10 | [8 empty] | 0.02553% | 72.29 | 66.57 | 3938.2812 | 3937.6492 | 0.2818% | 72.24 | 72.52 | 4290.2832 | 4292.5637 | 0.07975% | 72.39 | 72.63 | 4296.7908 | 4290.1667
avg monitor | 72.2784 | [8 empty] | 72.312 | 72.27 | 66.54 | 3936.5064 | 3936.9639 | 72.26 | 72.23 | 72.52 | 4290.2832 | 4293.158 | 72.338 | 72.39 | 72.52 | 4290.2832 | 4283.6691
|  |  | -26.9 | 18 | 0.99975 |  |  |  | 0.48971 | 2203.0198 | 72.26 | 37.25 | 2203.71 | 2204.2711 | 2310.2006 | 72.23 | 39.04 | 2309.6064 | 2311.154 | 2309.0335 | 72.28 | 39.12 | 2314.3392 | 2314.288
[10 empty] | 0.01926% | 72.3 | 37.23 | 2202.5268 | 2201.8688 | 0.1328% | 72.42 | 39.1 | 2313.156 | 2308.6332 | 0.0742% | 72.34 | 39.04 | 2309.6064 | 2307.6397
[10 empty] | 0.01546% | 72.19 | 37.2 | 2200.752 | 2203.4469 | 0.05608% | 72.37 | 39.07 | 2311.3812 | 2308.4556 | 0.1125% | 72.28 | 39 | 2307.24 | 2307.1889
[10 empty] | 0.01051% | 72.26 | 37.23 | 2202.5268 | 2203.0876 | 0.09489% | 72.19 | 39.06 | 2310.7896 | 2313.6193 | 0.1423% | 72.41 | 39.05 | 2310.198 | 2305.9994
[10 empty] | 72.27 | 72.34 | 37.26 | 2204.3016 | 2202.4246 | 72.308 | 72.33 | 39.06 | 2310.7896 | 2309.1411 | 72.326 | 72.32 | 39.07 | 2311.3812 | 2310.0516
|  |  | -36.6 | 22 | 0.99955 |  |  |  | 0.23309 | 1050.3447 | 72.23 | 17.73 | 1048.9068 | 1049.6097 | 1075.5038 | 72.33 | 18.23 | 1078.4868 | 1077.7174 | 1074.5278 | 72.34 | 18.18 | 1075.5288 | 1074.613
|  |  |  | +schuim | [5 empty] | 0.02139% | 72.32 | 17.78 | 1051.8648 | 1051.2597 | 0.08835% | 72.27 | 18.15 | 1073.754 | 1073.8788 | 0.07147% | 72.33 | 18.16 | 1074.3456 | 1073.5792
[10 empty] | 0.02918% | 72.39 | 17.78 | 1051.8648 | 1050.2432 | 0.1584% | 72.38 | 18.18 | 1075.5288 | 1074.0191 | 0.04605% | 72.27 | 18.17 | 1074.9372 | 1075.0621
[10 empty] | 0.01531% | 72.37 | 17.77 | 1051.2732 | 1049.9426 | 0.1532% | 72.21 | 18.18 | 1075.5288 | 1076.5476 | 0.06229% | 72.22 | 18.16 | 1074.3456 | 1075.2144
[10 empty] | 72.326 | 72.32 | 17.77 | 1051.2732 | 1050.6685 | 72.296 | 72.29 | 18.18 | 1075.5288 | 1075.3562 | 72.298 | 72.33 | 18.17 | 1074.9372 | 1074.1703
0.99962
conclusions
The recombination cannot be measured reliably as low as -40 V. 100 Volts seems to be the limit, see linearity.
The polarity correction is , as expected, extremely low. To get more detailed figures, a still better controlled experiment is needed.

## recombination
date: | 2014-07-22 |  |  |  | 1.0225 | -5.209e-07 | 0.045236 |  | VDOSE 1 |  | -100 |  |  |  | -400
linac: | U9 |  | table [cm] | RW3 [cm] | ks | Boag-Mw/Ml | u | cpp [nC] | Qrel | mon [nC] | Stealth [μC] | Qrel [*10-5] | Qrel | mon [nC] | Stealth [μC] | Qrel [*10-5]
field size: | 25x25 |  | at head | 0 | 1.0218 |  |  | 3.75 | 22.381 | 54.7 | 17.63 | 22.3944 | 23.815 | 54.67 | 18.75 | 23.8301
chamber: | NE2571, sn 2498 | [7 empty] | 0.06514% | 54.66 | 17.61 | 22.3854 | 0.09789% | 54.78 | 18.78 | 23.8203
electrometer: | PTW Webline, SN176, -250 V | [7 empty] | 0.04748% | 54.76 | 17.63 | 22.3699 | 0.07907% | 54.67 | 18.74 | 23.8174
phantoms: | - | [7 empty] | 0.04154% | 54.71 | 17.62 | 22.3776 | 0.04791% | 54.76 | 18.76 | 23.8037
depth: | var | [7 empty] | 54.708 | 54.71 | 17.62 | 22.3776 | 54.728 | 54.76 | 18.76 | 23.8037
ssd: | var |  | 14.6 | 0 |  | temperature? |  | 2.022 | 9.732 | 69.77 | 9.764 | 9.7237 | 10.1084 | 69.6 | 10.11 | 10.0929
Stealth-device: | A | [7 empty] | 0.08075% | 69.66 | 9.753 | 9.7281 | 0.06049% | 69.56 | 10.12 | 10.1087
electrometer: | DOSE 1 | [7 empty] | 0.08275% | 69.79 | 9.774 | 9.7309 | 0.08269% | 69.5 | 10.11 | 10.1074
temperature: | - | °C | [6 empty] | 0.07552% | 69.71 | 9.766 | 9.7341 | 0.09597% | 69.57 | 10.13 | 10.1172
energy: | 6 MV | [7 empty] | 69.722 | 69.68 | 9.771 | 9.7433 | 69.548 | 69.51 | 10.12 | 10.116
dose: | 100 | ME | -7.6 | 0 | 1.0073 |  |  | 1.2984 | 6.3578 | 69.47 | 6.357 | 6.3581 | 6.4961 | 69.48 | 6.492 | 6.4922
gantry: | 180 | ° | [6 empty] | 0.06186% | 69.55 | 6.364 | 6.3578 | 0.1075% | 69.34 | 6.487 | 6.5003
[9 empty] | 0.06976% | 69.57 | 6.365 | 6.357 | 0.05554% | 69.39 | 6.491 | 6.4996
Vw/Vl | 4 | [7 empty] | 0.02103% | 69.5 | 6.358 | 6.3564 | 0.05861% | 69.45 | 6.493 | 6.496
avg monitor | 69.4823 | [7 empty] | 69.53 | 69.56 | 6.367 | 6.3599 | 69.438 | 69.53 | 6.497 | 6.4925
|  |  | 14.6 | 0 | 1.0117 |  |  | 2.032 | 9.8273 | 69.52 | 9.82 | 9.8147 | 10.1681 | 69.47 | 10.16 | 10.1618
t | 12.5 | s | [6 empty] | 0.2241% | 69.38 | 9.81 | 9.8245 | 0.1041% | 69.51 | 10.17 | 10.1659
pulses | 5000 | [7 empty] | 0.07413% | 69.19 | 9.82 | 9.8615 | 0.088% | 69.4 | 10.17 | 10.1821
[9 empty] | 0.1989% | 69.54 | 9.829 | 9.8208 | 0.08005% | 69.36 | 10.15 | 10.1679
[9 empty] | 69.438 | 69.56 | 9.826 | 9.815 | 69.454 | 69.53 | 10.17 | 10.163
|  |  | -24.6 | 0 | 1.0055 |  |  | 0.9782 | 4.8269 | 69.42 | 4.825 | 4.8293 | 4.9064 | 69.31 | 4.891 | 4.9032
[9 empty] | 0.04757% | 69.47 | 4.823 | 4.8239 | 0.05237% | 69.39 | 4.901 | 4.9075
[9 empty] | 0.03968% | 69.46 | 4.827 | 4.8286 | 0.08968% | 69.39 | 4.902 | 4.9085
[9 empty] | 0.05238% | 69.5 | 4.827 | 4.8258 | 0.04187% | 69.37 | 4.899 | 4.9069
[9 empty] | 69.4625 |  |  |  | 69.372 | 69.4 | 4.9 | 4.9058
|  |  | -60.6 | 0 | 1.003 |  |  | 0.5952 | 2.9473 | 69.6 | 2.95 | 2.945 | 2.9737 | 69.51 | 2.976 | 2.9748
[9 empty] | 0.09725% | 69.43 | 2.944 | 2.9462 | 0.07472% | 69.47 | 2.974 | 2.9745
[9 empty] | 0.1006% | 69.48 | 2.949 | 2.9491 | 0.04984% | 69.59 | 2.976 | 2.9714
[9 empty] | 0.06672% | 69.56 | 2.952 | 2.9487 | 0.05895% | 69.59 | 2.977 | 2.9724
[9 empty] | 69.512 | 69.49 | 2.948 |  | 69.54 | 69.54 | 2.978 | 2.9755
|  |  | -105 | 20 | 1.0002 |  |  | 0.19028 | 0.95026 | 69.37 | 0.9474 | 0.94893 | 0.95083 | 69.52 | 0.9514 | 0.95088
[9 empty] | 0.05995% | 69.43 | 0.9492 | 0.94992 | 0.1122% | 69.51 | 0.9515 | 0.95112
[9 empty] | 0.1134% | 69.4 | 0.9501 | 0.95123 | 0.1269% | 69.35 | 0.9491 | 0.95091
[9 empty] | 0.0926% | 69.4 | 0.9496 | 0.95073 | 0.02284% | 69.4 | 0.9495 | 0.95063
[9 empty] | 69.384 | 69.32 | 0.9483 | 0.95052 | 69.432 | 69.38 | 0.9492 | 0.9506
|  |  | -105 | 20 | 1.0007 |  |  | 0.1904 | 0.95044 | 69.47 | 0.9505 | 0.95067 | 0.95234 | 69.41 | 0.952 | 0.95299
[9 empty] | 0.1231% | 69.35 | 0.9505 | 0.95231 | 0.0785% | 69.44 | 0.9508 | 0.95138
[9 empty] | 0.04102% | 69.54 | 0.9509 | 0.95011 | 0.07166% | 69.42 | 0.9516 | 0.95245
[9 empty] | 0.1202% | 69.51 | 0.9499 | 0.94952 | 0.06641% | 69.36 | 0.9511 | 0.95278
[9 empty] | 69.488 | 69.57 | 0.9508 | 0.9496 | 69.428 | 69.51 | 0.9525 | 0.95212
conclusions
There is a strong but linear dependancy on the charge per pulse for the recombination.

## stability
2014-08-13
Nucletron HDR-192Ir-source at approximately 3 meter.
Webline @ +400 V |  | meter al vorige dag ingeschakeld met aangesloten kamer
kamer A
tmeting [s] | 60
T½ [d] | 73.829 |  |  | correctie
tijd | T [°C] | p [hPa] | Q [nC] | verval | pT | I [pA]
11:25 | 21.2 | 1003 | 1.218 | 1 | 1.0144 | 20.5914
12:40 | 21.1 | 1002 | 1.223 | 1.0005 | 1.015 | 20.6997
12:46 | 21.1 | 1002 | 1.215 | 1.0005 | 1.015 | 20.5651
12:55 | 21.1 | 1002 | 1.214 | 1.0006 | 1.015 | 20.5493
13:25 | 21.1 | 1002 | 1.215 | 1.0008 | 1.015 | 20.5703
13:27 | 21.1 | 1002 | 1.212 | 1.0008 | 1.015 | 20.5198
15:38 | 21.2 | 1002 | 1.213 | 1.0017 | 1.0154 | 20.5612
15:45 | 21.2 | 1002 | 1.208 | 1.0017 | 1.0154 | 20.4774
15:48 | 21.2 | 1002 | 1.207 | 1.0017 | 1.0154 | 20.4609
16:40 | 21.2 | 1002 | 1.208 | 1.0021 | 1.0154 | 20.4847
16:45 | 21.2 | 1002 | 1.205 | 1.0021 | 1.0154 | 20.4345
18:05 | 21.2 | 1003 | 1.208 | 1.0026 | 1.0144 | 20.4757
18:08 | 21.2 | 1003 | 1.204 | 1.0026 | 1.0144 | 20.4083 | drift: | -0.14574 | [%/h]
18:12 | 21.2 | 1003 | 1.204 | 1.0027 | 1.0144 | 20.4088
18:15 | 21.2 | 1003 | 1.203 | 1.0027 | 1.0144 | 20.3922
18:18 | 21.2 | 1003 | 1.205 | 1.0027 | 1.0144 | 20.4266